## Slide 1: CEPC HCAL Readout Electronics

## Slide 2: Content
Background
  CEPC Calorimeters
  CEPC HCAL
  USTC SDHCAL
  SDHCAL Readout ASIC
Design of the Test-Board
  Introduction of Microroc
  Design of the Phase 1 Test-Board
  Progress
Future Work
  Full Test of Detector
  Design a Read-out Array

## Slide 3: Content
Background
  CEPC Calorimeters
  CEPC HCAL
  USTC SDHCAL
  SDHCAL Readout ASIC
Design of the Test-Board
  Introduction of Readout ASIC
  Design of the Phase 1 Test-Board
  Progress
Future Work
  Full Test of Detector
  Design a Read-out Array

## Slide 4: CEPC Calorimeters
Key Calorimeters
  Electromagnetic calorimeter (ECAL)
  Hadron calorimeter (HCAL)
Designed for precise energy measurements of electrons, photons, taus and hadronic jets.
Basic Resolution Requirements
  ECAL: 16%/√(𝐸(𝐺𝑒𝑉))
  HCAL: 50%/√(𝐸(𝐺𝑒𝑉)).
To achieve these, a Particle Algorithm oriented calorimetry system is being considered as the baseline design.

## Slide 5: CEPC Calorimeters
Particle Flow Algorithm (PFA)
  Charged particles –>tracker
  Photons – >electromagnetic calorimeter
  Neutral hadrons –> hadronic calorimeter
  Approximately 72% of the jet energy is measured
Both ECAL and HCAL require high granularity

## Slide 6: CEPC HCAL
HCAL
  A high-granularity HCAL plays an essential role in CEPC. It allows separation of energy deposits from charged and neutral hadrons. The measurement accuracy of the neutral hadrons is the leading contribution to the jet energy resolution for jets with energy up to 100GeV.
  The HCAL is a sampling calorimeter with steel as the absorber and scintillator tiles or gaseous devices with embedded electronics.

## Slide 7: CEPC HCAL
HCAL Detector Options
HCAL
AHCAL
DHCAL
SDHCAL
Scintillator + SiPM + SPIROC(ASIC), 8148channels,1m3
Analog Hadron Calorimeters(AHCAL).
High power consumption, expensive readout electronics
GEM/RPC/Micromegas + DCAL(ASIC), chn/m3
A hadron calorimeter with only one threshold readout is called a Digital Hadron Calorimeter (DHCAL).
Simple, but poor accuracy. Hard to satisfy PFA
GEM/RPC/Micromegas + HARDROC(ASIC), chn/m3
A more general calorimeter with multi-threshold readout (e.g. 3 thresholds) is therefore also considered, a so-called Semi-Digital Hadron Calorimeter (SDHCAL).
A balanced choice for CEPC HCAL
Absorber: steel
Active layers: Gaseous detector

## Slide 8: USTC SDHCAL
Architectural of the SDHCAL
A scheme of the ASU used to readout 1m2 detector
?
cm
,
)
Absorber(Fe, W)
Sensor(RPC,MM,GEMTHGEM)
PCB

## Slide 9: SDHCAL Readout ASIC
Hardroc (HAdronic Rpc Detector Read-Out Chip)
Microroc (MICRO mesh gaseous structure Read-Out Chip)
Comparison
Parameters | Hardroc | Microroc
Dynamic range | 10fC ~ 10pC | 2fC ~ 500fC
Channel per ASIC | 64 | 64
Shaper | 3 fast shapers per channel | 2 adjustable shapers per channel
Threshold | 3 thresholds | 3 thresholds
Power consumption | 10μW/Channel | ≈0 (Standby mode, power-pulsed)

## Slide 10: SDHCAL Readout ASIC
We choose GEM detector as our phase-I detector. Considering the dynamic range and other factors, the proper ASIC for GEM is Microroc

## Slide 11: Content
Background
  CEPC Calorimeters
  CEPC HCAL
  USTC SDHCAL
  SDHCAL Readout ASIC
Design of the Test-Board
  Introduction of Microroc
  Read-Out Structure for Future
  Design of the Phase I Prototype
  Progress
Future Work
  Full Test of Detector
  Design a Read-out Array

## Slide 12: Introduction of Microroc
Over view
  Analog Part
    160 pin QFP package or bare die
    Low noise PAC: < 2fC
    2 adjustable shaper: high gain and low gain
    3 discriminators and triple 10-bit DAC for threshold
    4 bit-DAC for base line adjustment
  Analog to Digital Interface
    RS to latch the trigger
    Single channel data convert to 2-bits.
  Digital Part
    Data length for one fire is 160-bits(Including header and the Brunch ID)
    160bits*128 RAM
    Open collector buffer
    Redundancy of the OC lines

## Slide 13: Analog Part of Microroc
The noise of the PAC have been simulated and found equal to 0.25 fC which allows to set the threshold of the discriminator that follow the HG shaper to a very low value such as 2fC.
1fc~500fC
Base line adjust
Low Gain
High Gain

## Slide 14: Digital Part
Single Channel Encode
RAM Data Structure
  Max event:127
<Discri2> | <Discri1> | <Discri0> | Encode<1> | Encode<0>
0 | 0 | 0 | 0 | 0
0 | 0 | 1 | 0 | 1
0 | 1 | 1 | 1 | 0
1 | 1 | 1 | 1 | 1

## Slide 15: Digital Part
Redundancy of OC Line
  To ensure the reliability the Doutb, transmitonb, start and end readout signals have been doubled
Daisy Chain
  The ASIC can work in daisy chain mode. Thus in ASU, the ASIC can be arranged in a matrix. Each row lay out in daisy chain and each column in parallel
Power Consumption
  Max power consumption: 10μW/Channel ->640μW/ASIC
  The power pulsing mode consists in switching off the bias current of the various cell in interbunch
  With power pulsing and 0.5% duty-cycle, mean power consumption is 3.2μW/ASIC

## Slide 16: SDHCAL Readout for Future

## Slide 17: ECAL&HCAL Readout For Future

## Slide 18: Phase I Prototype
Structure of Prototype
The connector among boards is 4-layers soft PCB
All the boards are completed

## Slide 19: Phase I Prototype
SDHCAL DAQ Board

## Slide 20: Phase I Prototype
Microroc Test Board

## Slide 21: Phase I Prototype
GEM Read-Out Pad Array

## Slide 22: Progress
FPGA Logic Completed
DAQ Software Completed (C# WPF Program)

## Slide 23: Progress
Single ASIC Electronics Test Completed
  Linearity of triple 10-bits DAC
  Uniform of shaper output base line
  S-Curve Test—”Trigger-efficiency vs Threshold”
  Linearity between charge and shaper output
  Hit Map of the Detector

## Slide 24: Linearity of 10-bit DAC
10-bit DAC
  The 3 thresholds of Microroc is set by triple 10 bit DAC
  The DAC code can be set by DAQ software via the SC parameters
  The DAC output voltage is calculate by the formula:
  , thus
  If a 200kΩ resistance is added between pin 73 and v_bg, =>

## Slide 25: Linearity of 10-bit DAC
Result
  We test two kinds of DAC. One has the 200k resistance between pin 73 and v_bg and the other does not.
  The second result is optimim
Slope:2.43 mV/DAC Unit
Max INL: 1.0869%
Slope:1.37 mV/DAC Unit
Max INL: 0.0776%

## Slide 26: Uniform of Base Line
Base Line
  The shaper output is negative, thus the base line has an offset voltage. In Microroc the offset is 2.2V set by the ASIC.
  There are discrepancies among 64 channels, thus it is essential to uniformize the offset voltage
Method to Uniformize
  The ASIC has a 4-bits DAC for each channel to adjust to high gain shaper output offset.
  The 4-bits DAC Code can be set by DAQ software via SC parameters.
  The offset voltage can be calculate as fallow:

## Slide 27: Uniform of Base Line
Result
  After correction, the standard deviation of the offset voltage is 0.3215mV.

## Slide 28: S-Curve Test
“Trigger-efficiency vs Threshold” Test.
  We define Trigger-efficiency as:
  Use a signal generator to inject the charge and the FPGA counts the trigger and total charge at the same time.
  Sweep the threshold(10-bit DAC output) from minimal to maximal, and get a curve call S-Curve.
S-Curve
  The ideal S-Curve in our system is a fermi distribution.
  Due to the noise of the circuit, the S-Curve has dispersion.
  With a certain input, the DAC output corresponding to 50% trigger-efficiency is the value of the shaper output

## Slide 29: S-Curve Test
Result Without Input
  Bigger DAC code means lower threshold. So when DAC code grows, the trigger-efficiency become higher.

## Slide 30: S-Curve Test
Small Charge Inject Test(< 10fC)
  Choose the 50% trigger-efficiency DAC code as the threshold, and get the linear fit of charge and DAC Code
  1fC Charge is corresponding to 3.9992DAC unit

## Slide 31: S-Curve Test
Noise of Electronics
  Statistics of none input 50% trigger-efficiency DAC code
  The standard deviation is 0.6326DAC unit corresponding to 0.1581fC

## Slide 32: S-Curve Test
Linearity of Two Shapers
  The linear range of high gain shaper is 0~140fC
  The linear range of low gain shaper is 0~500fC

## Slide 33: Hit Map
We use 8keV X-Ray to test our DAQ system and get the hit map

## Slide 34: Content
Background
  CEPC Calorimeters
  CEPC HCAL
  USTC SDHCAL
  SDHCAL Readout ASIC
Design of the Test-Board
  Introduction of Microroc
  Design of the Phase 1 Test-Board
  Progress
Future Work
  Full Test of Detector
  Design a Read-out Array

## Slide 35: Full Test of Detector
We have measure the electronics performance of the Microrc and our test-board
The next step is to combine the Microrc with the GEM to get full test.
We will test follow parameters
  Noise of base line
  Homogeneity of read-out pads
  Crosstalk between read-out pads
  Detection efficiency of GEM

## Slide 36: Design a Read-out Array
When all the test done, our next step is to design the ASU array.
ASIC in each row controlled in daisy chain mode and each column controlled in parallel